## Slide 1: Proline Promass

## Slide 2: Measuring principle

## Slide 3: Why Measure Mass?
Temperature / Pressure
Volume 1  Volume 2
Mass 1 = Mass 2
Masse 1
Volumen 1
Masse 2
Volumen 2
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 4: Traditional Mass Flow Measurement
Discontinuous, direct
m = Smi
Continuous, indirect
•
•
m = V • r
•
V = Volume flow
r = Density
m = Mass flow
•
mi
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 5: Measurement Signal
w=Angular velocity
Fc=Coriolis force
=time shift
A,B=Sensors
y=Amplitude
t=Time
 ~ Fc ~ m
•
If fluid is flowing, the Coriolis Force introduces a time shift of the swinging points A and B. Point B passes zero before point A . The bigger the flow the bigger the shift.
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Play Movie

## Slide 6: Frequency-shift by Mass-shift
Mass:
10 kg
Spring constant: 20 N/m
Amplitude: 10 cm
Mass:
1 kg
Spring constant: 20 N/m
Amplitude: 10 cm
Oscillating time: 4.4s
Max. velocity: 0.20 m/s
Oscillating time: 1.40 s
Max. velocity: 2.00 m/s
The resonance frequency of a swinging system is dependent on the stiffness c and the mass m.
The bigger the mass at constant stiffness, the lower the resonance frequency
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 7: The higher the density the bigger the mass
mt
fR
mfl
fR = Resonant frequency
mt = Tube mass
mfl = Fluid mass
fl = Fluid density
c = Constant
fR = ƒ(fl)
DN 25
Density
Air
Water
950 Hz
740 Hz
1,3
1000
The higher the density, the lower the resonance frequency
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 8: Overview of direct measuring variables
D=Phase shift
m=Mass flow
fR=Resonance frequency
r=Density
W= Resistance (PT1000)
T=Temperature
 ~ m
fR ~ 
 ~ T
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 9: Overview of calculated values
V=Volume flow V=m/r
VN=Normvolume flow = Volume flow at fixed p and T VN= m/rN (note: rN is a fixed value for each fluid)
c=Concentration Concentration can be calculated from density  see specific training (advanced module)
n=Viscosity Viscosity can be calculated from oscillation damping. Viscosity measurement is only available with the Promass I sensor
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 10: So far, the best coriolis flowmeter on the market
Not affected by :
  temperature change
  pressure change
  external vibration
Tested and calibrated in metrological laboratories complying to ISO 17025
AS GOOD AS A SWISS WATCH
All
our sensors and electronics are produced in Switzerland, with the precision of a Swiss time piece
Notes: The coriolis flowmeter is one of the state of the art technology and behind the principle there are 30 years experience on the field and improvement over the last decades. We invest a lot of money on the research and development in order to meet the fast moving and demanding expectations. And to prevent from temperature and pressure, vibrations influence, we develop a well balanced system using the best material spiced up with calibration know how. The advantage of such principle operate independantly of temperature, pressure, viscosity, conductivity, flow profil.
When you want to invest on a good quality Cronograph, you will go for Swiss, if you want a silk tie, you will go for french brand, for robust and quality shoes you will go for italian shoes? Won´t you? Therefore if you want a Coriolis flowmeter go for Endress+Hauser.
Is a manufacturer really capable of delivering high precision equipment ? Do instruments maintain their once proven accuracy over time even under changing conditions.
It consists of two variables : 1) the intitial accuracy which is verified by calibration 2) The long term accuracy which is determined by the design.
A watch is only a piece of jewelry. But in a business a wristwatch is considered as a personal statement of taste and status symbol. A perfect watch should be exclusive and yet modest according to my provider like Rolex, Breitling, Omega! And of course precision is absolutely outstanding even under harsh environement. And like our Coriolis, several test are made to test the reliability and the performance
( temperature test, shock test 200 times at 100G) and in the end it has to meet some criteria, for instance in clockwork a maximum permissible daily deviation of +/- 0,07 s that represent 2.10^-4.
What about calibration?
- Comparison of the measurement results of one instrument to the results of the reference instrument. An accredited calibration laboratory has to prove that its reference has a significantly better accuracy than the calibrated meters.
E+H flowtec has been accredited by SAS as a calibration laboratory for flowmeters as per the ISO norms like the independent laboratories for clockwork so called COSC has accredited its services by SAS as per ISO.

## Slide 11: Flow accuracy identical for volume and mass
 | Mass flow Uncertainty (liquid) | Uncertainty of volume flow (liquid)
Promass 83F PremiumCal | 0,05% | ---
Promass 83F | 0,10% | 0,10%
Promass 83S | 0,10% | 0,10%
Promass 83P | 0,10% | 0,10%
Promass 83I | 0,10% | 0,10%
Promass 83A | 0,10% | 0,10%
Promass 83H | 0,10% | 0,10%
Promass 83E | 0,25% | 0,25%
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 11
what you have to fathom here is same uncertainty either for volume and massflow according the flow sensors. F sensors is dual tube and most accurate and I sensors is single straight tube.

## Slide 12: Linear accuracy
Linear accuracy from 0,5 m/s
Linear accuracy Highest turndown 25:1
New
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 12
The accuracy that we specify is valid almost within the whole measuring range from 0,5 m/s
The highest turndown in the market 1:25

## Slide 13: Standard density calibration
Calibration with air and water under reference conditions.
Standard Certificate
Accuracy : ± 0.5 kg/m3 at ± 50kg/m3 and ±10°C of reference conditions.
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 13

## Slide 14: Special density calibration
Guaranteed accuracy within the whole working conditions of the device
No adjustment required in situ
Linearized device in terms of temperature and density
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 14

## Slide 15: Special density calibration
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 15
If the option special density calibration is selected the Promass will be calibrated with different fluids (=densities) and the temperature characteristics of the device is determined with several measurements at different temperatures with water. This procedure allows to specify a much higher accuracy that is valid within a range of 5…80°C / 1…2000 kg/m3. The densities of the reference fluids range from 820 to 1800 kg/m3.
Most applications are actually within the specified range and a written certificate allows a traceable statement of the performance of an individual meter. This is very unique since no other manufacturer is capable to provide a similar document. That leaves their customers to believe what the supplier promises in his specs. Worse still most suppliers make a statement for reference conditions or for field calibration whilst there is NO accuracy specified for any other (=actual process) conditions.

## Slide 16: Density specification
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 16
Standard : 2 points Air et eau
Précision : ± 0.5 kg/m3 pour masse volumique à ± 50kg/m3 et ±10°C des conditions de références.

## Slide 17: Segmentation

## Slide 18: Promass Sensors
Promass A
Promass E
Promass F
Promass M
Promass I
Promass H
Promass F HT
Promass P
Promass S
PHASE OUT
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 18

## Slide 19: Product Portfolio of Proline Promass
Single Tube Sensors for special applications
Dual Tube Sensors for highest accuracy requirements
The Economic Sensor for replacing volumetric flow-meters
Performance
80
83
A, I, H, S, P
M, F,
40
E
Sensors
Transmitters
Basic Standard High End

## Slide 20: The Coriolis Range
The most complete range
9 sensor models (DN1…350)
0,08 kg/h to 4100 tons/h
Stainless steel 904L, 316L, Alloy C-22, titanium and zirconium
Working Conditions
  -200°C to 350  C
  Up to 400 bar
Performance uncertainties :
  Lin. accuracy > 0,5 m/s
  Mass flow up to 0,05 %
  Density : ± 0, 5 kg/m3
  Accuracy volume = masse
New
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 20

## Slide 21: Proline 40: The basic Economical solution
Communication
0/4..20mA; 0...1kHz
HART
Service interface (Connection to Tooling)
Up to 3 cable glands
Human Machine Interface (HMI)
Blind version
optionally: 2-line back-
lit display in plain text
Measuring values
m, V
Data storage
S-DAT™
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Digital Com

## Slide 22: Transmitters

## Slide 23: Promass 200 - 2 wire : The basic Economical solution
Communication
2 wires
0/4..20mA; 0...1kHz
HART
Service interface Connection to Tooling)
Up to 3 cable glands
Human Machine Interface (HMI)
Blind version
optionally: 4-line back-lit LCD display in plain text
Extended diagnostics NE107
Measuring values
m, V
Data storage
S-DAT™, datalogging, liner recorder
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
SIL

## Slide 24: Promass 80: The Standard Transmitter
Communication
0/4..20mA; 0...1kHz
HART + Profibus PA
Service interface (Connection to Tooling)
Up to 5 cable glands
Human Machine Interface (HMI)
Quick Setup for a fast and easy commissioning
Push-button operation
2-line back-lit display: Display in plain text
Measuring values
m, V, , T
Data storage
S-DAT™
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 25: Promass 83: The High-End Transmitter
Communication
Flexible I/O-Modules
HART, PA, DP, FF, Modbus, Ethernet
Service interface (connection to Tooling)
Human Machine Interface (HMI)
Quick Setup for a fast and easy commissioning
Touch Control operation allows programming without opening of electronics
4-line back-lit display
Modular Software concept
Concentration calculation (Brix, Plato, °API, etc)
Batching functions
Viscosity measurement and
Advanced diagnostics
Data storage
S-DAT™, T-DAT™, F-CHIP™
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Digital Com

## Slide 26: Sensors

## Slide 27: Promass E200: Features
Flow Measuring Uncertainty:
Liquid Mass ± 0.5%
Liquid volume ± 0.5%
Gas Mass ± 1%
Flow Measuring Repeatability:
Mass/volume Flow (liquid): 0.25%
Mass (gas) : 0,5 %
Flow Range Liquid
DN8 2000 kg/h
DN15 6,500 kg/h
DN25 18,000 kg/h
DN40 45,000 kg/h
DN50 70,000 kg/h
DN 80 180, 000 kg/h
Available with Ex-Certificates:
ATEX
FM/CSA
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Plan Safety

## Slide 28: Promass 82BE: Features
Flow Measuring Uncertainty:
Liquid Mass ± 0.25%
Liquid volume ± 0.25%
Flow Measuring Repeatability:
Mass/volume Flow (liquid): 0.25%
Mass (gas) : 0,5 %
Density uncertainty:
0,0005 kg/l
Flow Range Liquid
DN8 2000 kg/h
DN15 6,500 kg/h
DN25 18,000 kg/h
DN40 45,000 kg/h
DN50 70,000 kg/h
Certificates:
ATEX
FM/CSA
SIL proven in use
NE107 Namur for diagnostics
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Plan Safety

## Slide 29: Promass 80/83 E: Features
Flow Measuring Accuracy
Liquid Mass : ± 0.30/0.25%
Gas Mass : ± 0.75%
Flow Measuring Repeatability:
Mass Flow: 0.01%
Flow Range Liquid
DN8 2000 kg/h
DN15 6,500 kg/h
DN25 18,000 kg/h
DN40 45,000 kg/h
DN50 70,000 kg/h
DN80 180,000 kg/h
Density Measuring Uncertainty:
Field Calibration: ±0.0005 kg/l
Entire range/sensors: ±0.020 kg/l
Temperature Measuring Uncertainty: 0.5° ± 0.005 x T
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 30: Promass E Sensor Construction
Simple and robust design.
Sensor housing may contain pressure up to 15 bar
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 31: Promass 80/83 F: Universal Solution
Application
Low pressure gases (e.g. natural gas, air)
LPG (e.g. butane, propane)
Fuel oils and gasoline
Fruit concentrates
Vegetable oils, animal fats
Alcohol
Detergents and solvents
Latex, Silicon oil
Toluene, Benzene
Cryogenic gases to -196°C
“The standard for almost
every application, our most accurate meter!”
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 32: Promass 80/83 F: Features
Flow Measuring Uncertainty:
Liq. Mass : ± 0.15/0,1% ± 0.05% optional for 83F
Gas Mass : ± 0.35%
Flow Measuring Repeatability:
Mass Flow: 0.05%
Flow Range:
DN8 2000 kg/h
DN15 6,500 kg/h
DN25 18,000 kg/h
DN40 45,000 kg/h
DN50 70,000 kg/h
DN80 180,000 kg/h
DN100 350,000 kg/h
DN150 800,000 kg/h
DN250 2,200,000 kg/h
DN 350 4,100,000 kg/h
Density Measuring Uncertainty:
Field/Ref. Calibration: ±0.0005 kg/l
Special Calibration: ±0.001 kg/l
Standard Calibration: ±0.010 kg/l
Repeatability: 0.00025 kg/l
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 33: Promass F: Sensor Design
Electrodynamic sensors
Exciter
Secondary containment rated up to 40 bar
Measuring tube
Rigid carrier tube
The rigid carrier tube withstands all loads in a piping system, allowing for true “Fit and Forget”.
No extra measures like pipe line support needed.
No restrictions on pipe mismatch.
Filled with Nitrogen
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 34: Also available in high temperature version
Technical data
Fluids: Liquids and gases
Nominal diameter: DN25 (1”), DN50(2”), DN80(3”)
Temperature: -50...350°C
Tube Material: Alloy C-22
Flow splitter material: SS 316L (optional Alloy C-22)
Flange material: SS 316L (optional Alloy C-22)
Connections: Flanges DIN PN40,
ANSI Cl.150, 300
JIS 10K, 20K
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 35: Promass 80/83 A: When Every Drop Counts
Applications
High-pressure gases (up to 400bar)
Low pressure gases (e.g. helium, hydrogen, air)
Additives and flavors
Perfume
DI-water
Insulin
Glue
“Sometimes the little things in life
count the most!”
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 36: Promass 80/83 A: Features
Flow Measuring Uncertainty:
Liquid Mass : ± 0.15/0,1%
Gas Mass : ± 0.50%
Flow Measuring Repeatability:
Mass Flow: 0.05%
Flow Range
DN1 20 kg/h
DN2 100kg/h
DN4 450 kg/h
Density Measuring Uncertainty:
Field/Ref. Calibration: ±0.001 kg/l
Special Calibration: ±0.002 kg/l
Standard Calibration: ±0.020 kg/l
Temperature Measuring Uncertainty:
  ± 0,5°C
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 37: Promass A: Sensor Design
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 38: Promass 80/83 I: Features and Benefits
Diameter range 3/8"...3” (DN 8 ... 80)
Single straight tube design, Full-bore version
Low pressure drop reduces the cost of ownership.
"Gentle" treatment of measured liquid.
Fully drainable.
Torsion mode balanced system TMB™
No pipe line support needed = Lower installation cost
extended vibration immunity.
Fully welded version
No gaskets: = reduction of maintenance cost
Wetted parts from Titanium
High corrosion resistance
Can be used up to 150°C
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 39: Promass 80/83 I: The Specialist
Application
Yogurt with fruit pieces
Syrup/Molasses
Chocolate with nut pieces
Blood plasma (sterile)
Detergents, solvents
Colors
Liquefied gases
Cosmetics
Also measures viscosity
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 40: Promass I Patented Viscosity control
Diameter DN 8 ... 80
Wide range of application and industries.
Low pressure drop
No shearing stress
Single straight tube
Perfect for waste oil metering
Torsion Mode Balancing TMB™
Dual frequency oscillation means the measurement is immune to bubbles etc
No fixing accessories required
Lower installation cost and better vibration immunity
Fully welded no gaskets
Reduction of possible leakage points and maintenance cost
...with flow rates 10kg/h...180’000 kg/h
Notes: The torsional movement of the measuring tube can be used to measure the viscosity of the fluid.
Due to the rotational motion v of the tube, the fluid is forced to a rotational motion. Depending on the viscosity, we find different velocity profiles v of the
fluid. Due to the velocity gradient of the profile there are shear forces in the fluid.
These shear forces damp the measuring tube and can be measured via excitation current, which is necessary to maintain the tube oscillation.
By measuring the excitation current (tube damping), the viscosity of the fluid can be determined

## Slide 41: Promass I Construction
Fixation 2
Exciter
Fixation 1
Measuring tube
Torsion bar
Sensor 1
Sensor 2
Balancing tube
Exciter et eccentric sensors/ balancing tube. Optimized for a better stability on inhomogeneous fluids and viscosity measurement
Performance : liquid flow 0, 1% of mV + zero stability (density 2 kg/m3)
Gas flow 0,5% of mV + zero stability

## Slide 42: Internal view
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 43: Optimized balancing system
DUAL MODE Promass I is the first Coriolis using two simultaneously resonance frequencies. This means measurement of viscosity and mass flow without interference.
Lateral mode 600-700 Hz
  Massflow
  Density
Torsion mode (900Hz)
  Viscosity
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: Do Promass I meters still use TMB balancing?
TMB comprises the overall know-how of E+H single tube technology. According to TMB vibration energy is transformed into the torsional motion which can be decoupled from environment. The balancing concept of Promass I redesign can be conceived as a novel implementation of TMB.
Due to different resonance frequencies and the newly introduced dual mode the results for viscosity measurements in non-newtonian fluids may differ.
It is now possible to order the viscosity software package without special density calibration!
Massflow and zero point specifications have been improved significantly.
Stability and performance of mass flow and viscosity measurements at multiphase conditions have been improved significantly
Air bubbles in the process result in an increased exciter current due to the dampening effect. For the initial Promass I this affected also the viscosity measurement and resulted in unstable performance.
eddy current brake has been removed from the internal structure. This improves the performance at inhomogeneous process conditions dramatically.

## Slide 44: Accuracy and measuring ranges
Accuracy with newtonian fluids: ± 5% ± 0.5 mPas (cP)
Repeatability: 0.5%
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 45: Promass S, the Sanitary meter for the Food Industry
Mass flow (liquid): ± 0.1% (83) ± 0.15% (80)
Volume flow (liquid): ± 0.25% (83) ± 0.3% (80)
Standard density: ±0.01 g/cc
Special density: ±0.002 g/cc
Field density / calibration or under reference conditions: ±0.0005 g/cc Repeatability = 0.00025g/cc
Temperature: - 50°C to 150°C
There is no better single tube coriolis meter with stainless steel tubing!
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 46: Measuring tube material
Promass S
Measuring tube 1.4539 / 904L for outstanding corrosion resistance.
Process connection 1.4435 / 316L
Surface finish 0.8um
EHEDG certified, 3-A approved,
pigable, for cost saving cleaning
The meter that meets all the requirements of the food industry
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 47: Promass S hygienic process connections
DIN 11864-2 Form A
DIN 11864-3 Form A
ISO 2853
DIN 11864-1 Form A
ISO 2852
DIN 32676
DIN 11851
Tri-Clamp
SMS 1145
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 48: Temperature shocking, CIP/SIP simulation
To ensure the flow meter can handle thermal shock as experienced in the CIP process.
Temp cycled between 10°C and 140°C
No effect found on meter performance afterwards
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 49: Promass Portfolio: Certification and Hygienic Design
Dosimass
Promass A
Promass E
Promass S
Promass F
Promass I
Temp >135 °C
3-A
EHEDG
Custody Transfer
Piggable
Drainable



















Certification
Hygienic Design


Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 50: Promass 80/83P the Pharma meter
Promass P
Certification for highly regulated industries: ASME BPE Certificate of Compliance to relevant scope
Inspection certificates: EN 10204 3.1, MTR; - for material, surface roughness, and delta ferrite.
Material selected according to ASME BPE - 1.4435 / 316L, low delta ferrite
Wetted surfaces: Ra max=0.76 µm or Ra max=0.38 µm
Electropolished flow tube and process connection and electropolished exterior: for sterile, easier cleanability and higher corrosion resistant surface
Hygienic fully welded, sensor design for full drain ability
“Designed for CIP/SIP” Continuous measurement with fast recovery - no delay
Accredited flow calibration
  according to ISO/IEC 17025 (SCS/A2LA),
  traceable density calibration
Comes with all the paper work so you can focus on the “not so compliant” parts of your plant
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 51: Measuring tube material
Promass P
Measuring tube 1.4435 / 316L (low delta ferrite)
Process connection 1.4435 / 316L (low delta ferrite)
Surface finish Ra max = 0.76µm (mechanical polish)
Surface finish Ra max = 0.38µm (electropolished)
EHEDG certified, 3-A approved
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 52: Promass P Measuring Accuracy
Mass flow (liquid): ± 0.1% (83) ± 0.15% (80)
Mass flow (gas): ±0.50%
Volume flow (liquid): ± 0.25% (83) ± 0.3% (80)
Standard density: ±0.01 g/cc
Special density: ±0.002 g/cc
Field density / calibration or under reference conditions: ±0.0005 g/cc
Temperature: - 50°C to 200°C
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 53
Flow Measuring Uncertainty:
Liquid Mass/Vol: ± 0,1/0.15% ± zero point
Gas Mass : ± 0.50% ± zero point
Flow Measuring Repeatability:
Mass Flow: 0.05% ± 1/2 zero point
Flow Range
DN8 2000 kg/h
DN15 6,500 kg/h
DN25 18,000 kg/h
DN40 45,000 kg/h
DN50 70000 Kg/h
Density Measuring Uncertainty:
Field/Ref. Calibration: ±0.0005 kg/l
Special Calibration: ±0.002 kg/l
Standard Calibration: ±0.020 kg/l
Temperature Measuring Uncertainty:
± 0,5°C
Promass 80/83 H: Features
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 54
Secondary containment
Electrodynamic sensors
Exciter
Measuring tube
Balancing Weight
Balancing Tube
Balancing Tube Weights
Measuring and Balancing tube couplers
The Endress+Hauser patented ITB – System (Intrinsic Tube Balancing):
An attached balancing tube and Balancing weights ensure perfect balancing allowing for true “fit and forget”
Slide 54
Endress+Hauser Consult AG Sales Support
Promass H: Cross Sectional View
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 55: Application for tantalum
Tantalum has the widest spectrum in corrosive environment
Zirconium become an alternative in reducing area
Corrosion test by CIBA (chemical company)
  Tube + flanges face + welding
  HCL 32% @ 95°C
  H2SO4 96% @120°C
Important reminder:
  Chemical compatibility has to be validated by the customer !
  Samples are available on requirement
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: For the time being, we do never specify a life time for a sensor in specific media or apllications. Since the life time in aggressive media can be heavily influenced by changes in temperature, concentration or contaminants!
55
18.10.17

## Slide 56: Some examples of applications for tantalum
Typical applications – Masse Flow / Volume flow, density and concentration measurement :
  32% HCL (hydrochloric acid) @ 150 °C
  96% H2SO4 (Sulfuric acid) @ 120 °C
  65% nitric acid @ 100 °C
  85% phosphoric acid @ 100 °C
Concentration measurement of acids:
  Sulfuric acid @ 1840 kg/m3: ±1% of concentration  ±8 kg/m3 in density
  Hydrochloric acid @ 1500 kg/m3: ±1% of concentration  ±5 kg/m3 in density
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 56
18.10.17

## Slide 57: Advantages of Promass

## Slide 58: Features / Promass Endress+Hauser
Compact
  Most compact design, lightweight
Secondary containment (excepted sensors E)
  Robust construction, improved process security
Immune to vibrations
  No affected by pipe vibrations
Self drainable
  No dead volume, available as hygienic version
Performance
  0.05% with linear precision (>0.5 m/s)
  The least affected by temperature or pressure
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 58

## Slide 59: Compact Construction
Comparison of sensors 3" (DN 80) Promass
Lightweight
Required little place
No need for support or brackets
Simple Installation
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 59
Customers require ease of installation in all three dimensions, not only short face to face length. Important are also ease of access in order to wire the meter.
Lighter weight means also the possibility to install without mechanical support of the pipeline.
Quality requirements in some cases may mean to remove the meter for verification or recalibration, meaning ease of handling is important.
Example comparison in the picture:
Promass M DN80 / 3”
Length = 840 mm
Height = 385 mm
Weight = 67 kg
Emerson MicroMotion
CMF300
Length = 856 mm
Height = 1071.5 mm
Weight = 77 kg

## Slide 60: Standard Secondary containment (excepted sensor E)
Serve as security if the fluid is extremely corrosiv or abrasiv
Specifications are checked by burst test for each diameter (Test certificate delivered by a third part)
Enable to optimized sensors stability for high accuracy insitu
Secondary containment
Measuring tube
Nitrogen
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 60

## Slide 61: Promass: No sensitive to pipe vibrations
Balanced measuring system
High working frequency ( typically 600 to 1000 Hz)  no interference with oscillations from pipe vibrations (normally around 50 à 250 Hz)
Barreau vibrant
Vibration effect from competition sensors
Vibration effect from Endress+Hauser promass
No twisting occurs on critical points
= Perfect protection
Notes: 61
All Coriolis mass flowmeters are based upon the same principle. The measurement of the mass flow can be achieved with only one measurement tube. But by using only one measurement tube (and no intelligent compensation system), the system is more sensitive against vibration. That is why most manufacturer are offering mass flowmeter with 2 measurement tubes. The 2 tubes are vibrating in opposite direction, making the system an intrinsically balanced one.
This means that no vibration from the vibrating tubes will go outside. Also vibration from the outside is less likely to affect the vibration of the measurement tubes, unless the outside vibration has the same frequency as the meter.
For Promass, the oscillation frequency of the measurement tubes is very high, between 200 and 1000 Hz, typically around 400 to 600 Hz. Meters from the competition usually oscillate between 50 and 200 Hz. In this frequency region, there usually are pipeline vibrations caused by pumps or compressors. If this pipeline vibration is around the same frequency as the Coriolis meter oscillation frequency, the meter can be affected. For this reason competition meter usually are more sensitive to vibration.

## Slide 62: Self drainable construction
For all sensors
  No dead volume
  No settlings problem
  No corrosion
  No sensitive to coatings
  Certified EHEDG, 3-A
Self drainable
With air and liquid
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 62
Each of the Promass sensors is fully drainable if installed in a vertical position. Designs with loops or other designs as found in the portfolio of some competitors don’t offer this advantage.
Depending on the design of the Promass model part drainage of the sensor is often possible sometimes only with a minimal slope. In particular the P sensor is offered with eccentric Clamps that allows to FULLY drain the device even in a horizontal mounting position. This is a requirement often found in the pharmaceutical industry. Promass P is the only Coriolis flowmeter offering this advantage.

## Slide 63: Performance - Coriolis E+H
Accuracy : 0.30% à 0.05%
Linear precision for v > 0.5 m/s
Temperature effect 0.0002%°C
Negligeable pressure effect
Special density calibration possible
The most accurate calibration rig in the world 0.015%, traceable according to ISO17025
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 63

## Slide 64: Coriolis Endress+Hauser
High accuracy, zero stability, little sensitive to process variations (P, T°)
Lightweight sensors, compact and robust
Drainable sensors from top to bottom
No sensitive to vibrations
No influence linked to viscosity variations
No inlet or outlet runs required
No filters required
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: 64

## Slide 65: Coriolis flowmeters
Installation guidelines

## Slide 66: Coriolis installation considerations: no straight inlet run
Coriolis meters are not affected by flow profile.
Advantages for installation:
  Coriolis flowmeters do not require straight inlet or outlet runs.
  Elbows, valves, or pumps upstream do not affect the performance of the meter.
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: Most other modern flow principles are based on measuring the fluid velocity and calculate flow from this measurement. In contrast, Coriolis measurement is a direct measurement of mass flow. As a consequence, Coriolis meters do not require straight inlet or outlet runs which are necessary for most other measurement principles. This can be a big advantage when pipeline length is restricted or for the construction of a skid.
For comparison an example of straight run requirements for other measurement technologies after a 90º elbow:
Magmeters: 5xDN straight inlet, 2xDN straight outlet
Ultrasonics (92F): 5xDN straight inlet, 5xDN straight outlet
Vortex meters: 20xDN straight inlet, 5xDN straight outlet
Thermal meters: 15xDN straight inlet, 5xDN straight outlet

## Slide 67: Coriolis installation considerations: full pipe (1)
The measuring tube(s) of a Coriolis meter must be kept full.
Promass triggers “Empty pipe detection” (EPD) warning if density falls below defined limit.
Consider the following recommendations:
  Best mounting location: In vertical pipeline with flow streaming upwards
  Good alternative: Horizontal orientation on a low point in the pipeline.
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: The best mounting location for a Coriolis meter is in a vertical pipeline with the flow streaming upwards. With this orientation, the pipe is either full or completely empty and the empty pipe detection will respond fast and correctly.
If mounting on a vertical pipeline is not possible, a good alternative is installation in a horizontal pipeline on a low point in the pipeline system.
The empty pipe detection (EPD) is a Promass software function which triggers a warning if either the density falls outside limit values which can be chosen according to the application. Factory setting for the low limit values is 0,2 kg/l. If in a liquid application the pipeline is emptied, the density will drop below that limit value and the EPD function will trigger a warning which can be displayed and/or signalled via the status output.

## Slide 68: Coriolis installation considerations: full pipe (2)
Consider the following restrictions:
  Avoid installation at the highest point of a pipeline.
  Avoid installation directly upstream of a free pipe outlet in a vertical pipeline.
If installation before a free pipe outlet is inevitable:
  Decrease nominal outlet diameter (Pipe restrictor or an orifice plate).
  Consult operating manual for recommended orifice sizes.
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: Mounting at the highest point in the pipeline has to be avoided in any case because in this position the pipeline can easily be partially filled, thus impairing the measuring accuracy of the Coriolis meter.
Installation in a vertical pipeline with downwards flow has to be avoided as well. In this case, the back pressure of the Coriolis meter is equal the (ambient) outlet pressure which can result in the sensor running empty during measurement.
However, there can be situations when this is the only possible installation point for the meter (e.g. in the Chemical or Life Sciences industry when educts are batched into the reactor). In these cases it is recommendable to use pipe restrictors or orifice plates with smaller cross section downstream of the meter. This measure increases the back pressure of the meter which ensures a full measuring tube and thus accurate measurement.
The Promass operating manual lists the recommended inner diameter of a pipe restriction or orifice plate, depending on the nominal diameter of the meter in question.

## Slide 69: Coriolis installation considerations: entrained solid
With curved measuring tubes and horizontal installation:
Entrained solids in liquids: do not install with curves pointing downwards (risk of solids accumulating).
Steam measurement: do not install with curves pointing downwards (risk of condensate accumulating).
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: The two cases below have a similar symptom and the same solution yet the root cause is different.
Solids
If solids are present in a liquid there is a risk that the solids settle at the bottom of the tube if the flow is stopped. Worst case the solids “bake” to the tube wall and build up over time. Any build up on the tube wall can cause an unbalance that can lead to an unstable measuring signal and measuring errors.
Condensate in gas/steam
Generally liquids in a gas measurement are not acceptable since two phase fluids are difficult or even impossible to measure accurately. But even in applications where this is not a general problem it might be that any present condensate would collect in the Promass if it is installed in a horizontal pipe. If this occurs the same influence as above will result.
The problems can be minimized or avoided if the sensor is installed under an angle or/and a model is selected that is fully drainable (Promass I or Promass P with eccentric Clamps)

## Slide 70: Air entrainment
All (Coriolis) flowmeters are affected by entrained air.
Promass in combination with FieldCare can be used to detect unknown gas entrainment problems in application (by tracking the stability of the meter’s excitation current).
Effects on measurement depend on the air distribution in the fluid.
Bubbles
Slug flow
Partial filling
Or any kind of combination
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: All flowmeters, including all Coriolis meters, are affected by entrained gas. Promass meters do not stall measurement with entrained air; however measurement accuracy will be impaired. The effect on meter performance depends on air amount and air distribution in the liquid.
If there is entrained air in the liquid, the damping of the measuring tube will increase. In consequence the excitation current needed to drive the oscillation of the measuring tube will be higher than in homogeneous flow conditions. This effect can be used diagnose entrained air. In the Promass software, a warning on display and/or output can be generated when the excitation current exceeds a definable maximum value. In combination with the FieldCare operation software, the stability of the meter’s excitation current can be tracked over a longer time period which allows more accurate and more convenient detection of air entrainment.

## Slide 71: Possible air entrainment effects on measurement
Bubbles
Noisy erratic reading depending
on bubble size and quantity
Slug flow
Delayed response mechanism,
erratic reading depending on slug size and frequency. Periodic noise depending on slug and pipe size.
Partial filling
Extremely noisy reading, huge errors depending on amount of air, instrument should switch to error mode. Product splashes through the pipes.
Any combination?
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 72: Coriolis installation considerations: entrained air
Do not install with curves pointing upwards (risk of gas accumulating).
If entrained gas (e.g. viscous fluids) is present, increase back pressure (to about 6-10 bar) to dissolve the gas in the liquid and improve performance.
Install on the high-pressure side of a pump and as close as possible to the pump.
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: In an application with unwished air entrainment, real improvement can only be achieved by removing the air. The most promising measure is always to increase the application pressure (6-10 bar). Installation on the high-pressure side on the pump and as close as possible to the pump can help to achieve this goal.

## Slide 73: Coriolis installation considerations: zero point stability
Zero point stability is defined by design and manufacturing tolerances of the flowmeter.
Promass normally does not need to be zeroed because of its exceptional zero point stability. (Still it is often common practice to perform a zero point adjustment.)
Occasional zero point adjustment can be considered under the following application conditions (if the flow rate displayed is stable but has an incorrect value):
  High viscosity or inhomogeneous fluid
  Possibility of abrasion or corrosion of the flowmeter
  Mechanical stress on the flowmeter
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Notes: NOTE:
Process induced zero point fluctuations might not be stable and constantly the same. Therefore any zero point adjustment is only a temporary remedy removing influences that are present at the moment.

## Slide 74: Installation hints
Promass F
With gas content
Promass F
With solids content
Avoid installation on suction side of pumps => outgasing of liquids
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 75: Installation hints
Only use one fixed support
To avoid cavitation extend reduced length
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications

## Slide 76
90° Bend
Flow devider:
Heavy particles collect on the outer curve, lighter one on the inner one.
Solutions:
Turbulences to get a homogeneous mixture
Flow conditioner
Increase the system pressure = use valve or orifice after meter
Gas separator before the meter
Installation hints- liquids with gas content
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications
Commissioning

## Slide 77: Heating Jackets available, order structur DK8HJ-XXXXX
Heating Jacket mounting bolts
Laser “Button” weld
Flow diversion weld
Heating fluid process connection
Promass F Sensor
Available for:
Promass A
Promass E
Promass F
Promass I
Promass M
Keeps your process warm
Principle
Segmentation
Transmitters
Sensors
Advantages
Installation
Accessories
Applications